Portable monitor for registration of ballistocardiographic signals

Masloboev Y.P.

Moscow State Institute of Electronic Technology (Technical University)

Ballistocardiography (BCG) is a method by which body movements caused by heart activity are registered. BCG provides a potential application to monitor heart condition. The main advantage of this method is that there is no electrodes required to be attached to the body during the measurements.

The ballistocardiographic system as a whole consists of the cardiovascular system as a primary source of force, of a human body acting as a connecting link and of the ballistocardiographic monitor. We constructed a portable monitor for registration and processing of BCG signal. This device can be used as wearable system for monitoring of heart health in zero-gravity conditions of space flight.

Monitor consists of accelerometer, analog-digital converter (ADC) and software for signal processing. Accelerometer consists of steel disk with piezoceramic structure. The copper load is attached to disc. For analog-digital converting Sigma-Delta ADC is used.  Monitor includes internal memory for storing of BCG signal for 8 hours.

The main technical parameters:

- acceleration range 0.001 – 0.1 m/c2 ;

- frequency range 0 – 260Hz;

- weight – 120g;

The software for signal processing is based on the artificial neural network algorithms and can be used for:

- extraction and analysis of periodical components of signal like heart and respiratory rhythms;

- extraction and analysis of nonperiodical components of signal like movement activity and artifacts.