Real-time shaping of the embedded waveform for external defibrillation

Sergey V. Selishchev

Department of Biomedical Systems, Moscow State Institute of Electronic Technology (Technical University)

Objectives

The question of what current waveform should be selected as optimum, which independent from transchest resistance dynamics needs further elucidation. Two opportunities: forming anyone waveform, and providing this one independently from tranchest resistance dynamics – are needed for experiments.

The paper presents a method and apparatus for shaping the electrical waveform delivered by an external defibrillator based on an electrical parameter measured during delivery of the waveform. Software and hardware co-design allow embedding anyone waveform and realizing it independent from transthoracic electrical impedance changes during shock time.

Methods

The both method and apparatus are based on developed high voltage delivery unit.

The power cell block (PCB) delivers of high voltage to the patient through low pass filter. PCB consists of energy storage capacitors and insulated-gate bipolar transistors under pulse-width modulation (no more than 50 kHz) with low-stress switching for efficiency.

Voltage sensor is measuring PCB voltage output and current sensor is measuring PCB current output. This data transfer to input of the PCB control circuit through scaling transducers.  The value of scaling transducers depends upon the setting defibrillation energy.

The setting power waveform stores into power etalon waveform generator (PEWG). This etalon may be anyone. Its power waveform is embedded by software technology. Besides the integral from power waveform etalon on shock time is setting defibrillation energy.

The PCB control circuit is real-time driving the process of high voltage delivery using real-time comparison of PEWG output and scaling transducers output, that is control feedback.

The high voltage delivery unit was simulated by PSpice. Than it was realized by printed-circuit board technology.

Results

There are illustrations of simulation results for high voltage delivery to patient. The simulation model includes the skin breakdown with 100V threshold. There is example of asymmetric biphasic waveform etalon, which now is optimum for external defibrillation. The both duration the first and second phases are 4 ms at the half-level amplitude maximum of each phase. The amplitude of second (inverted) phase is half of first phase, and pulse tilt is 30 %.

Figures illustrate the voltage-time curves for 200 J delivered energy and 150, 50, 25  load resistance, and for 5 J, 25 . This simulation results correspond to the worst combination for electrical parameters of the circuit components.

The experimental results are presented too.

Conclusions

There are both a method and apparatus for shaping the electrical waveform delivered by an external defibrillator based on the PCB output measured and control feedback. Software and hardware co-design allow embedding anyone waveform and realizing it independently from transthoracic electrical impedance changes during shock time.